## LSR
LSR II | SN: H47200101 | BD: (877) 232-8995
Date | Case # | Problem | Solution
2022-05-02 | 1674113 | Arm to the side, dripping | flushed back side of DCM line
2023-09-22 | 2138258 | low FCS/ SSC, scattered CST, taking sample (not clogged) dripping
2023-09-29 | 2138258 | No samples, beads and cleaning solutions taking up, looks clogged
2023-10-26 |  | standby takes sample, not through flow cell, | air pressure valve
2024-08-19 |  | Red laser had no signal and can't see QC beads on Red laser detectors | Removed the old red laser and Installed the used red laser on 10/20/2024

## Canto
Canto II | SN: V96100817 | BD: (877) 232-8995
Date | Case # | Problem
2022-07-25 | 1750673 | Status error preventing Start-Up ("System prime fail")
2022-09-08 | 1790683 | fluid
2023-01-03 | 1895527 | Status error preventing Start-Up ("System prime fail")
2023-01-20 | 1910466 | fuidics, high currents, emergency fluidic shut down
2023-03-03 | 01951998 | fluidics not connected
2023-05-10 | 2014721 | tube getting stuck, hoder not letting go
2023-05-11 | 2015871 | not sucking up the flush
2023-12-21 |  | PM
2023-09-22 |  | sheath waste low error and requested shut down
2024-03-20 | 02300179 | BLUE LASER HIGH CURRENT
2024-04-24 | 2332705 | failed fluidc start up
2024-06-10 |  | PM
2024-06-28 | 2395645 | Waste sensor not working (William Crowel)
2024-08-08 | 2435210 | hardware issue
2024-08-22 |  | installed win 7 computer and Dive 8 software
2024-08-30 |  | Jeff comes to make cusomized configuration
2024-12-05 |  | PM and Flow Cell Replacement
2025-01-29 | 2630485 | system prime failed, float switch timeout. build up by drain | prime after tank refill x2. take top caff off of filter on right of the wet cart. purge bubble filter
2025-03-06 | 2675389 | weird sound from wet cart | replaced compressor
2025-06-02 |  | PM SCHEDULED
2025-12-18 |  | PMI Service Scheduled | PMI Service was done
4/092026 | 3061087 | burning plastic, cart | Vaccum pressure and burning issues resolved
2026-05-29 |  | PMI Service Scheduled | PMI Service was done

## Aria
Aria II | SN: P18100008 | BD: (877) 232-8995 | Comments
Date | Case # | Problem | Solutions
2022-08-29 | 1780532 | can not turn on L/R stream. unable to get sheath level, unable to turn stream on, hardware not responding
2022-09-06 | 1788277 | Waste sensor will not connect to fluidics cart. When waste container tried on Canto fluidics, it connected
2022-11-02 | 1843998 | wide blue laser peak, FSC-A is wavy, event rate is not stable for beads
2023-06-26 | 2057007
2023-09-05 | 2122563 | sheath sensor issue
2023-09-13 | 2130097 | low CST and accudrop events
2023-11-02 | 2176788 | Stream signal lost, FSC/SSC very low, Bead event rate is low
2023-11-22 | 2192872 | FSC/SSC very low
2023-12-04 | 2203997 | Side stream intensity low/can't sort Accudrop beads
2023-12-07 | 2207995 | Leaking even after Turn off stream (valve 6 behind nozzle area)
2024-01-31 | 2253999 | Compressor chamber pressure low 55psi
2024-04-04 |  | PM
2024-04-30 | 2337876 | loud air noise (talked with steve)
2024-07-24 | 2419929 | streams stuck in plate mode | (off and on over night resolved issue)
2024-09-23 | 2480409 | Fluidics cart leak from P3 pump | Replace pump
2024-09-24 |  | PM and pump repair
2025-03-31 | 2698065 | stream was not stable= pressure issue. Manually adjusted pressure for now | New Valve Stack + ordered new nozzles
2025-07-16 | 2799373 | stage thing keeps going up and stopping before it reaches sample line. it come back down. Error: Clog Switch Time out | Resolved the issues by Engineer
2025-09-15 |  | PM Service, Need to ckeck Accudrop for 85uM nozzle. | PM service done by Brad
2025-12-16 | 2956287 | Shaky Camera causes high amplitude, unstable stream & frequence change, Bend sample line causes clogg | Service was done and all looked good
2026-03-11 | 3033222 | Sample is not running on Flow Rate 1 and no agitation of sample tube, PMI service due | PM and other service was done 03-26 & 31-2026

## Aurora 3
Aurora | SN: R0026/CC291 | Cytek: (510) 657-0102 | Computer Name: SOMRADONCOL0045
Date | Case # | Problem | Solution
2022-0513 |  | small particle setting, noise 3000/s | Issue was solved
2022-1122 |  | red laser no signal | Replaced Red Laser
2023-0920 |  | Red laser negative or no signal | Red laser replacement
2023-1201 |  | QC Failed few times | Issue was resolved
2024-0605 | 340906 | Loud grinding noise when rinsing and mixing between the samples | Resolved by adjusting the up/down motor alignment
2024-0626 | 358739 | Loud grinding noise when rinsing and mixing between the samples | Resolved by adjusting the alignment which reduced the noise but parts need to be replaced
2024-04-17-18 |  | PMI Service and Software Upgrade | PMI and software upgrade was done
2024-0925 |  | PMI Service by Enginner | PMI Service was done
2024-0926 |  | Violoet Laser had signal issues and need replacement | Violet Laser was replaced
2025-0109 |  | Engineer Work on Violer laser mirror | Resolved the issue
2025-0220 | 572844 (cancelled) | Cyotometer doesn't take up volume of sample and looks sample line clogged | Sample Line (Yellow) was replaced ourselves as per Cytek Reco.
2025-04-29 | 560527 | PMI Service by Engineer work order #00478737 / Sample Line Struck occationally | PMI Service was done
2025-10-15 |  | PMI Service by Enginner on 10/20/2025 | PM service was done on 10/20/2025
2026-02-10 |  | PMI service and lousd noise from the Vacuum pumps | PM and other service was done on 03/27/2026
2026-06-11 |  | Cytometer not connecting and Flow meter died | Flow meter replaced and resolved the issues on 06/11/2026

## Aurora 4
Aurora-UV | SN: V0333 | Cytek: (510) 657-0102 | Computer Name: radoncsomcytek
Date | Case # | Problem | Solution
2022-0725 |  | Red Laser Neg Signal | Red Laser was replaced
2023-0125 |  | dripping and low side scatter odd sound | Resolved the issues
2023-0210 | 88958 | side scatter moving around every few samples | Resolved the issues
2023-0228 |  | side scatter moving and dripping | Resolved the issues
2024-0201 |  | Makes too much noise before/after running samples | Resolved the noise issues
2024-0429 | 280345 | LOUD noise around the sample probe area | issue was resolved by pulling whole sample stage down,
2024-0626 | 358343 | Replacement of APD detector array | APD detector was replaced
2024-09-26 |  | PMI Servie by Engineer | PMI Service was done
2025-04-29 | 560528 | PMI Service by Engineer work order #00478738 | PMI service done, Replaced UV Telescope & Dichroic Adjuster
2025-07-24 | 680722 | QC Failed Multiple times and UV detectors rCV was high | Service done and QC passed
2025-09-02 | 707951 | CPU power supply issues, likely replacement | Power Supply device was replaced on 09/10/2025
2025-10-15 |  | PMI Servie by Engineer on 10/20/2025, Pump down | PM service was done on 10/20/2025, pump fixed by Tim on 10/23/2025
2025-01-30 | 811617 | UV detector gains now approaching twice original value
2026-02-10 |  | PMI service and Flow cell re-coupling | PM and other service was done on 03/17-24/2026

## Aurora 5
Aurora-5 | SN: U1368 | Cytek: (510) 657-0102 | Computer Name: ???
Date | Case # | Problem | Solution
2023-1201 |  | QC Failed few times | Issue was resolved
2024-0401 | 295502 | R3 rCV to high, QC fail | Issue was resolved
2024-0402 | 296482 | probe does not come down in plate mode & makes sound during SIT calibration | Issue was resolved
2024-12-04 | 506870 | PMI Service and take up volume is high | PMI service was done and replaced Filter
2025-08-21 |  | PMI Service Scheduled by Cytek | PMI service was done and replaced Filter on 08/29/2025

## Aurora CS
Aurora CS | SN: S0205 | Cytek: (510) 657-0102 | Contract renewed until 06/28/2026
Date | Case # | Problem | Solution
2023-10-25 | WO 00179317 | Can't do QC and Beads/samples not running on Flow rate 1 | Resolved the issues by Engineer
2023-11-13 | 196940 | Can't do QC and Beads/sample LINE LEAKING | Resolved the issues
2023-12-14 |  | QC failed, sample line clog or not running on Flow rate 1-3 | Resolved the issues by Engineer
2024-02-01 |  | QC failed, Sheath line tube is leaking | Sample line changed, upgraded and Issue was resolved
02/26/2023-03/01/2024 |  | Aurora CS upgrade hardware and software | upgraded and all the issue was resolved
2024-06-26 | 358392 | 85 and 130 Nozzle installation and upgrades | Nozzle Installation was done and upgraded
2024-09-10 | 439396 | QC failed, FSC beads wider, PMI Service | PMI Service was done
2024-11-11 | 488696 | Slow dripping at the right side bottom under Qick stop button | Changed the female connector by ourselves
01/14-15/2025 | 525399 | PMI Service by Enginner (Work order #00449874) | PMI Service was done
2025-05-13 | 634542 | DDU Connection Instrument Error (Work order #00536413) | Sorter working after few days and cancelled the service
2025-06-17 |  | Sample was not running and less events | Changed Sample Line by ourselves
2025-07-28 | 675880 | PMI service request, Live camera disconnection, plate loader magnet makes sound | PM, camera disconnection was serviced 08/13-15/2025
2025-10-23 |  | Plate Loder needs to be fixed | Fixed plate loader
2026-05-11 | 882199 | Software crash and overfill below waste | Overfill waste leak was fixed
2026-05-11 | 793570 | PMI service, YG laser, Pressure issues, software closing during shutdown | PMI done & YG laser (06/8-12)and Pressure Regulator were replaced (06/22)

## MACSQuant
MACSQuant | SN: 21424 | Miltenyi: 866-811-4466
Date | Case # | Problem
2022-0718 |  | leaking and PM, cleaned inside of minisampler
2022-1230 |  | no events detected.
2024-1108
2025-0130 |  | PM
2025-08-25 |  | PM Service
2025-09-15

## FlowSight
FlowSight | SN: FS184 | Luminex: (877)785-2323 | support@luminexcorp.com
Date | Case # | Problem
2025-09-16 |  | PM Service scheduled on 10/15/2025 by Cytek | PM service was done by Sam / Tam

## Mili-Q
Mili-Q | SN: F8BN07863K | Milipore: (888)645-5478
Date | Case # | Problem
2024-0201 | C33234432 | need to replace Qpod